## Slide 1
Introduction
Methods
Results
The SUMMIT Study: Quality of spirometry performed as part of a ‘Lung Health Check’ (LHC) assessment for inclusion into low dose computed tomography (LDCT) lung cancer screening (LCS)
Sophie Tisi1, Jennifer L Dickson1, Carolyn Horst1, Helen Hall1, Priyam Verghese1, Andrew Creamer1, Anne-Marie Mullin2, Rachael Sarpong2, Jonathan Teague2, Laura Farrelly2, Kylie Gyertson3, Vicky Bowyer3, Claire Levermore3, Allan Hackshaw2, John R Hurst1, Sam M Janes1.1 Lungs for Living Research Centre, University College London, London, UK; 2 CRUK Cancer Trials Centre, University College London, London, UK; 3 University College London NHS Foundation Trust London UK
Discussion
Chronic Obstructive Pulmonary Disease (COPD) is substantially under-diagnosed, with rates of underdiagnosis estimated at 65-80%1. Incorporation of spirometry into LCS programmes may allow an opportunity for case-finding for undiagnosed COPD.
Spirometry has not been widely utilised within LCS to-date, with implementation of differing delivery protocols and no clear consensus as to how spirometry should be performed within this setting.
Poorly performed case-finding spirometry may cause harm through the identification of false positives, leading to unnecessary investigations, increased patient anxiety and overburden of healthcare services.
To ensure good quality, clinically useful spirometry within LCS, a quality assured approach to performance must be undertaken. There is recognition that quality assurance (QA) of all aspects of LCS must be undertaken but no guidelines exist around spirometry QA in this setting.
We report the SUMMIT Study’s (a study to assess the implementation of LDCT LCS and to validate a multi-cancer early detection blood test) approach to spirometry performance, quality assurance and quality assessment.
SUMMIT opened to recruitment in April 2019 in London. 55-77 year-olds with a GP coded smoking history were invited to a LHC where they undergo risk assessment for eligibility for LDCT, alongside a respiratory health consultation & pre-bronchodilator spirometry to screen for COPD.
Quality assured approach to spirometry
Equipment quality control
ATS/ERS equipment quality control criteria2 was ensured through the use of the Vitalograph Micro® handheld spirometer which meets ISO:26782 standards. Daily calibration was performed with a 3L syringe with a certified accuracy to within 0.5%. Full calibration and service logs were maintained.
2. Spirometry training & performance
A bespoke, in-depth training programme trained operators to ARTP standards which included direct observation of technique. Participants perform up to 3 manoeuvres and are coached to ensure acceptability in as many manouevres as possible.
Those with airflow obstruction (FEV1/FVC <0.7) who are symptomatic (cough >6 weeks and mMRC >1) and do not report a previous diagnosis of COPD are referred to primary care for consideration of diagnostic post-bronchodilator spirometry.
1000 spirometry sessions were assessed, with sessions performed by 24 operators (an average of 41.7 sessions per operator). 94.3% sessions contained 3 manoeuvres with a total of 2936 individual manoeuvres assessed. 73.6% (2160) manoeuvres met acceptability criteria, with a further 8% deemed usable, leading to 81.6% (2395) manoeuvres deemed either usable or acceptable.
At least 1 acceptable manoeuvre was seen in 91.4% (914) sessions, with 77.3% (773) sessions containing >2 acceptable manoeuvres and 3/3 manoeuvres meeting acceptability criteria in 46.5% (465) sessions. The commonest reason for manoeuvres not meeting acceptability criteria was a cough in the first second of expiration (seen in 10.8% manoeuvres) [Table 1].
Of sessions with >2 acceptable manoeuvres, 86% sessions met repeatability criteria of within 150ml for FEV1, 70% for FVC and 61.2% for both. 95.6% sessions had both FEV1 and FVC within 250ml [Table 1].
3. Central quality review.
The first 1000 spirometry sessions were reviewed centrally. Each manoeuvre was assessed against ATS/ERS within-manoeuvre criteria2 and graded as ‘acceptable’ ‘usable’ or ‘unacceptable’. Repeatability criteria were further assessed.
We report the development & implementation of an evidence-based approach to QA of spirometry performance within a LCS setting. We have demonstrated that good quality case-finding spirometry can be performed whilst maintaining a minimally invasive and time-sensitive approach. Our results reflect positively when compared to studies which have demonstrated poor quality of handheld spirometry within a community3. Full clinical utility of this approach will be assessed with longer-term diagnostic outcomes data.
References:
1.Diab N, Gershon AS, Sin DD, Tan WC, Bourbeau J, Boulet L-P, et al. CONCISE CLINICAL REVIEW Underdiagnosis and Overdiagnosis of Chronic Obstructive Pulmonary Disease. 2018 [cited 2021 Apr 16]; Available from: www.atsjournals.org. 2. Miller MR, Hankinson J, Brusasco V, Burgos F, Casaburi R, Coates A, et al. Standardisation of spirometry. 3. Ferguson GT, Enright PL, Buist AS, Higgins MW. Office spirometry for lung health assessment in adults: A consensus statement from the national lung health education program. Chest. 2000 Apr 1;117(4):1146–61.
Table 1: Percentage of manoeuvres meeting within-manoeuvre and between-manoeuvre criteria: